## Slide 1
ABSTRACT
RESULTS AND DISCUSSION
CONCLUSIONS AND FURTHER RESEARCH
[TEXT HERE]
ACKNOWLEDGMENTS
[TEXT HERE]
Breathe Providence: Towards a Distributed Method for Quantifying the Reliability of Low-Cost Air Quality Sensors
Mike Woodilla — Advisor: Professor Meredith Hastings
Low-cost sensors for measuring common air pollutants have made urban air quality research possible at high spatial and temporal resolution. This high-resolution data can be used to identify air quality events and trends at scales that are impractical for expensive reference-grade instruments to measure, however calibration of low-cost sensors remains challenging. Sensor applications in large distributed air quality monitoring networks further complicate calibration needs – laboratory calibration has been shown to translate poorly after sensor deployment and requires downtime for every sensor in the network, and in situ gas calibration requires frequent site visits to account for seasonal interferences to sensor measurements. Alphasense electrochemical gas sensors are also known to be cross-sensitive to temperature and relative humidity, which are seasonally-dependent and can cause poor sensor performance, especially in summer months, and induce drift over time.
This project seeks to characterize the reliability of carbon monoxide sensors in the Breathe Providence air quality monitoring network using distributed methods. Existing methods rely extensively on simultaneous reference observations and residual statistics, however availability of reference data from Environmental Protection Agency monitoring stations is limited in the City of Providence. Therefore, in addition to calculating residual statistics, the probability density function (PDF) of measurement and reference observations is approximated for each month since sensor deployment. When reference‑grade observations are not available, the first operating year of sensor data is treated as reference. F-Divergences of reference and measurement distributions are then calculated and interpreted as an indication of drift. Since low-cost gas sensors are thought to be unreliable at high temperature or relative humidity values, we compare residual correlation and PDF divergence with and without a meteorological filter which excludes observations at Temp > 30°C and RH > 0.75. This method is applied to a sensor with published reliability statistics as a control experiment.
To characterize the long-term deployment reliability of Breathe Providence Network CO and FPM sensors using existing frameworks.
To develop a new framework for sensor reliability characterization that minimally relies on external reference measurements.
To test this new framework against existing reliability characterizations.
To apply this new framework to the Breathe Providence Network and report findings.
As an experimental control, we first examine reliability statistics for BEACO2N @ Richmond Field Station (RFS), Berkeley CA; existing reliability statistics for this sensor are published by Winter et. al. (June 2025) which allows us to be confident in the true reliability state of the sensor. Time series correlation statistics are plotted over duration of deployment for the sensor. We find strong residual correlation with reference for this particular sensor over time, however we seek a reliability characterization method that relies minimally on simultaneous reference observations. We therefore estimate the probability density function (PDF) of sensor observations for each month of deployment and characterize the divergence between reference and measurement distributions using Kullback-Liebler, Hellinger, and Euclidean divergences. PDFs are estimated via a discrete histogram approximation with bin width set to 0.05 ppm; this approximation granularity was found to be necessary to reduce sensor noise uptake in the distribution estimation and to prevent floating-point arithmetic underflow. Since low-cost sensors are likely to be inaccurate at high temperature and relative humidity values, we compare residual correlation and PDF divergence with and without a meteorological filter which excludes observations at Temp > 30°C and RH > 0.75. We also compare trends in divergence when the first 12 months of the sensor’s operation are treated as reference distributions. We find that there is no significant trend in divergence when the source of reference data is either external or internal, nor do we find significant change in divergence before and after meteorological filtering.
Junior studying
Mathematics and
Computer Science
OBJECTIVES
Graphs of BEACO2N CO observation characteristics at Richmond Field Station, Berkeley CA, across months deployed.
Row 1: Time series residuals; Time series residual statistics; Probability density functions of observations from measurement and reference sensors;
Row 2: f‑Divergences with meteorological filtering; f‑Divergences without meteorological filtering; f‑Divergence change after meteorological filtering. Reference is external RFS Picarro CO sensor.
Row 3: Same as Row 2 but with distributions from first 12 operating months of BEACO2N sensor used as reference.
METHODS
Across the Breathe Providence network, we find that CO sensors do not exhibit significant trends in divergence from external or internal reference distributions with or without meteorological filtering; standard deviation of the measurement distribution exhibits seasonal variation but does not exhibit long term trends for any CO sensors in the network. This finding does not necessarily assert the reliability of our sensors, but merely indicates that no significant drift was detected in the network.
Furthermore, we find that meteorological filtering tends to increase distribution divergence for CO sensors in the Breathe Providence network as time since deployment increases. Of the [X] CO sensors examined, [X] exhibited a noticeable increase in KL divergence after meteorological filtering; this trend was exhibited both when external data and internal first-year data was used as reference. KL divergence has a standard interpretation as the information lost by using a measurement distribution to approximate a reference distribution, therefore an increase in KL divergence after meteorological filtering indicates that low-cost CO sensors lose information that is captured by reference when observations during strenuous meteorological conditions are omitted.
This finding indicates that more study is needed regarding how low-cost sensor observations at strenuous meteorological conditions are processed. To the author’s knowledge, no reliability or calibration studies of low-cost gas sensors in air quality monitoring applications specifically examine data recorded at high temperature or relative humidity values. Existing literature characterizes low-cost gas sensors as unreliable at these meteorological conditions, however this characterization is solely determined by examining correlation with reference instruments. Though low-cost gas sensors may not correlate one-to-one with reference-grade instruments at these meteorological conditions, this study suggests that these observations still carry information about the distribution of air quality during a given time period.
Alphasense CO-B4
Left: A deployed Berkeley Environmental Air-quality and Network (BEACO2N) sensor in the Breathe Providence Network.
Right: Internal components of a BEACO2N sensor, with the Alphasense CO-B4 sensor emphasized.